## ID NOW System
The following information is intended as an aid in the development of your facility's IQCP Risk Assessment and Quality Control Plan for the ID NOW™ test system. This document is not intended to replace the Package Insert or Instrument User Manual. Any modifications to this document are the sole responsibility of the facility.
All specific information here is taken from the ID NOW™ Influenza A & B 2 package insert IN427000 v. 5.0 2018/09 , ID NOW™ user manual INNAT024 v. 5.0 2017/09 and Technical Memorandum Amplicon Contamination Prevention SCRTECHS-0104 Rev. 1. Software 5.2.2 and the technical bulletin in TB000022 v10.0 2018/05. This is a waived complexity test, and IQCP is NOT required. This document has been developed for customers that requested a IQCP support document for the ID NOW™.
FAILURE | CAUSE | FAILURE TYPE | POTENTIAL EFFECT(S) OF FAILURE | ID NOW™ RISK MITIGATION FEATURES | ID NOW™ LABELING INSTRUCTIONS, WARNINGS AND PRECAUTIONS | LABORATORY RISK MITIGATION | SEVERITY | FREQUENCY | LABORATORY DOCUMENTATION
Operator failure | Not following manufacturer instructions | Testing Personnel |  | Guided instructions are on the screen of the instrument. | Package Insert (PI) and Instrument User Manual (UM)
First time instrument set-up | Wrong date set | Testing Personnel |  | Only Admin can set date and time displayed | Getting Started- Admin Setup section of UM
| Wrong time set | Testing Personnel |  | Only Admin can set date and time displayed | Getting Started- Admin Setup Menu section of UM
Instrument set up | Wrong power supply used | Testing Personnel |  | Power cord and adapters supplied with instrument. | Introduction/Instrument Specifications section of UM
| Wrong printer used | Testing Personnel |  | None | Accessories section of UM
| Wrong barcode scanner | Testing Personnel |  | None | Accessories section of UM
| Incorrect configuration | Testing Personnel |  | Instrument is configured according to facility’s needs with Admin feature | Getting Started- Initial Setup section of UM
| Incorrect user IDs and/or passwords set up | Testing Personnel |  | Admin can edit user ID/passwords | Getting Started- Admin Setup Section of UM
| First time QC not performed | Testing Personnel |  | When the instrument is used for the first time, the user is required to run both a positive and negative successful QC test before patient testing is allowed. | Running a Test section of UM.
Instrument environment | If too cold, elution buffer may not heat to temperature prior to start of test | Environment |  | None | Introduction/Instrument Specifications section of UM.
| Humidity out of range | Environment |  | None | Introduction/Instrument Specifications section of UM.
| Altitude and Pressure out of range | Environment |  | None | Introduction/Instrument Specifications section of UM.
| Light intensity | Environment |  | None | Introduction/Unpacking section of UM.
| Blockage of instrument air vents by walls, or other instruments | Environment |  | None | Introduction/Unpacking section of UM.
| Unstable surface | Environment |  | None | Introduction/Unpacking section of UM.
Inappropriate sample | Interfering substances with method | Specimen |  | None | Interfering Substances and Limitations section of PI
| Visibly bloody samples used | Specimen |  | None | Precautions section of PI
| Use of unapproved sample type | Testing Personnel |  | None | Specimen Collection and Handling Precautions, and Limitations section of PI
Incorrect sampling procedure | Sample contamination with other body fluids/materials during collection | Specimen |  | None | Specimen Collection and Handling, Precautions and Limitations sections of PI
| Incorrect sample type tested | Testing Personnel |  | None | Specimen Collection and Handling section of PI
| Incorrect technique for sample collection | Testing Personnel |  | None | Specimen Collection and Handling section of PI
| Use of kit control swab as patient swab | Testing Personnel |  | Instrument screen has clearly marked button for positive and negative controls. Swab packages are clearly labeled as positive or negative. | Unique Control swab labeling and packaging and Running a QC Test section of UM
| Use of non-validated swab | Testing Personnel |  | None | Specimen Collection and Handling section of PI
Incorrect sample handling | Contamination after sample collection | Testing Personnel |  | None | Specimen Collection and Handling section of PI, Introduction/Maintenance and Cleaning section of UM
| Incorrect sample temperature | Environment |  | None | Specimen Collection and Handling, Specimen Transport and Storage, and Limitations section of PI
| Swab not rotated in Transport Media for 10 seconds | Testing Personnel |  | None | Specimen Transport and Storage and Limitations section of PI
| Incorrect volume of Transport Media used | Testing Personnel |  | None | Specimen Transport and Storage section of PI
| Unsuitable Transport Media used | Testing Personnel |  | None | Specimen Collection and Handling, Specimen Transport and Storage, and Limitations section of PI
| Too long delay from collection to analysis | Testing Personnel |  | None | Specimen Collection and Handling, Specimen Transport and Storage, and Limitations section of PI
Operator failure | Not following PI and UM instructions | Testing Personnel |  | Guided instructions are on the screen of the instrument. | PI and UM
| Eluted swabs in Transport Media not mixed before testing | Testing Personnel |  | None | Specimen Transport and Storage and Limitations section of PI
| Inappropriate amount of sample presented from Sample Receiver to Test Base | Testing Personnel |  | Orange indicator does not fully rise, the Transfer Cartridge may not collect enough sample. | Test Procedure section of PI
| Incorrect patient identification information entered into instrument | Testing Personnel |  | Barcode scanner available. | Running a Test section of UM. Accessories section of UM
| Running a patient sample as a control or vice versa. | Testing Personnel |  | Manual selection required by Testing personnel on screen menu. | Running a Test section of UM
Instrument failure | Inadequate warm-up time | Test System |  | White light slow flash | Getting Started-Initial Setup/Instrument Status Light section of UM
| General Instrument failure - Barcode reader, touch screen, optic system, Test Base rotors etc. | Test System |  | Errors and Warnings codes | Errors and Warnings section of UM
Reagent handling | Storage of test components outside of specified range- Temperature | Environment |  | None | Storage and Stability section of the PI
| Storage of test components outside of specified range- Humidity | Environment |  | None | Introduction/Instrument Specifications section of UM
| Test Base pouch is received by user already opened | Reagent |  | None | Precautions section of PI
| User opens Test Base pouch too early and leaves it out | Testing Personnel |  | None | Precautions section of PI
| User uses scissors to open pouch and cuts PCR tube | Testing Personnel |  | None | Precautions section of PI
| Leaking or spilling Sample Receiver | Reagent |  | Sample Receiver is foil sealed. | Test Procedure section of PI
| Drops Sample Receiver and uses | Testing Personnel |  | None | Precautions section of PI
| Drops Transfer Cartridge and uses | Testing Personnel |  | None | Precautions section of PI
| Sample Receiver not at room temperature | Testing Personnel |  | None | Storage and Stability section of PI
| Different lots of reagents used | Testing Personnel |  | None | Precautions section of PI
| User reuses components after system has timed out or test cancelled | Testing Personnel |  | None | Precautions section of PI
| Outdated reagents | Testing Personnel |  | Expiration date is printed on reagents. | Precautions section of PI.
Reagent failure | Not following reagent handling instructions | Testing Personnel |  | Step by Step instructions shown on the instrument screen. | PI and UM
| General Reagent Failure- Amplification, Primer annealing, Nicking reaction, etc. | Reagent |  | ID NOW™ Influenza A & B 2 contains an internal control that has been designed to control for sample inhibition, amplification and assay reagent function. In positive samples where target amplification is strong, the internal control is ignored and the target amplification serves as the ‘control’ to confirm that the clinical sample was not inhibitory and that assay reagent performance was robust. At a very low frequency, clinical samples can contain inhibitors that may generate invalid results. | Quality Control Section of PI
| Reagent contamination | Testing Personnel |  | Procedural Control Valid displayed on the instrument screen indicates that the assay reagents maintained their functional integrity and the sample did not significantly inhibit assay performance. | Test Procedure section of PI
| Deterioration of reagent lots over time | Reagent |  | ID NOW™ Influenza A & B 2 contains an internal control that has been designed to control for sample inhibition, amplification and assay reagent function. In positive samples where target amplification is strong, the internal control is ignored and the target amplification serves as the ‘control’ to confirm that the clinical sample was not inhibitory and that assay reagent performance was robust. At a very low frequency, clinical samples can contain inhibitors that may generate invalid results. | Reagents must be stored according to the PI and UM.
Testing- loading consumables | Incorrect disposable insertion (Sample Receiver inserted into Test Base position and vice versa.) | Testing Personnel |  | Procedural Control Valid displayed on the instrument screen indicates that the assay reagents maintained their functional integrity and the sample did not significantly inhibit assay performance. | Test Procedure section of PI
| Wrong assay (i.e. different iNAT test) | Testing Personnel |  | Instrument reads barcode and instrument screen indicates test type and asks user to confirm | Test Procedure section of PI
| Mis-match of Sample Receiver and Test Base (i.e. Flu and Strep A) | Testing Personnel |  | None | Test name is on the package labeling and foil seal of the Sample Receiver.
| Barcode not recognized due to Test Base misalignment | Test System |  | Instrument performs a rotor alignment before each test | Test Procedure section of PI
| User damages barcode | Testing Personnel |  | Instrument prompts user to initiate a second read of barcode. | Errors and Warnings section of UM
| User waits more than ten minutes after inserting Test Base. | Testing Personnel |  | Instrument will timeout after 10 minutes. ("Error timeout" screen) | Test Procedure section of PI
| Incomplete insertion of Sample Receiver | Testing Personnel |  | Test will not run if Sample Receiver is not inserted correctly and cannot be recognized by the Instrument. | Test Procedure section of PI
| Sample Receiver is inserted at same time as Test Base | Testing Personnel |  | Instrument screen instructions. User can remove Sample Receiver and put back in to make instrument recognize it. | Test Procedure section of PI
| User walks away after insertion of Sample Receiver | Testing Personnel |  | Instrument screen instructions. Instrument will timeout after ten minutes. ("Error timeout" screen) | Test Procedure section of PI
| User empties elution buffer into Sample Receiver heating block | Testing Personnel |  | User can't transfer sample (won't aspirate) | Test Procedure section of PI
| User takes Sample Receiver out of instrument during warm up period | Testing Personnel |  | Instrument screen instructions to put Sample Receiver back in and starts a warm up again. | Test Procedure section of PI
| User takes Sample Receiver out of instrument to remove foil seal (after heating) | Testing Personnel |  | Instrument screen instructions say to put back in and starts a warm up again | Test Procedure section of PI
| User closes lid during warm up | Testing Personnel |  | Instrument screen instructions are on the screen of the instrument, and the message will not change if lid is closed during warm up. | Test Procedure section of PI
Testing- eluted sample/swab elution | Elution buffer spills during removal of foil seal. | Testing Personnel |  | Instrument holds Sample Receiver in place to some degree | Test Procedure section of PI
| Use of object to puncture seal | Testing Personnel |  | None | Test Procedure section of PI
| User does not take foil seal off, and pushes OK button | Testing Personnel |  | Instrument screen instructions provided when to remove the foil seal. | Test Procedure section of PI
| Wrong volume of eluted swab in transport media used | Testing Personnel |  | Instrument screen instructions provided to add 200 ul of sample | Test Procedure section of PI
| Eluted swab in transport media not mixed vigorously | Testing Personnel |  | Instrument screen instructions provided to mix sample for 10 seconds | Test Procedure section of PI
| Swab not eluted thoroughly | Testing Personnel |  | Instrument screen instructions state to mix swab for 10 seconds | Test Procedure section of PI
| Swab is not vigorously mixed. | Testing Personnel |  | Instrument screen instructions for mixing. | Test Procedure section of PI
| Swab/Sample is eluted prior to warm up | Testing Personnel |  | Instrument screen instructions for Sample Receiver warm up and sample addition. | Test Procedure section of PI
Sample Transfer | Delay in transferring the sample | Testing Personnel |  | Instrument times out after 10 minutes, instruction to dispose of components | Errors and Warnings section of UM
| User does not insert Transfer Cartridge properly into the Sample Receiver | Testing Personnel |  | Audible and tactile feedback; Orange plunger rises (visible from top of Transfer Cartridge) | Test Procedure section of PI
| Elution buffer (with a potential positive sample) drips from Transfer Cartridge onto instrument | Testing Personnel |  | Sample Receiver and Test Base protects sample from contaminated surfaces of instrument. | Test Procedure section of PI
| Delay in transfer to Test Base after aspirating sample. | Testing Personnel |  | Instrument times out after 10 minutes, instruction to dispose of components | Errors and Warnings section of UM
| User does not properly connect the Transfer Cartridge with the Test Base or press down to transfer the sample | Testing Personnel |  | Feedback (clicking and orange plunger descending), Instrument recognizes Transfer Cartridge position on Test Base | Test Procedure section of PI
| User continues to apply pressure to Transfer Cartridge after it's connected (more than 30 seconds) | Testing Personnel |  | Feedback features (clicking and orange plunger descending). | Test Procedure section of PI
| User applies too much pressure to Transfer Cartridge after it is connected. | Testing Personnel |  | Feedback features (clicking and orange plunger descending), seal of disposables strong enough to withstand this force, user manual | Test Procedure section of PI
| Use of Transfer Cartridge that has not aspirated sample. | Testing Personnel |  | Instrument screen instructions. Audible and tactile feedback; Orange plunger rises (visible from top of Transfer Cartridge) | Test Procedure section of PI
Amplification/Test results | Too long to close lid | Testing Personnel |  | Instrument screen instructions. Instrument times out after 30 seconds, instruction to dispose of components | Errors and Warnings section of UM and Test Procedure section of PI
| User opens lid during run | Testing Personnel |  | Instrument screen instructions state to not open lid. Test cancels if lid is opened during test run, and user is instructed to dispose of test materials. | Errors and Warnings section of UM and Test Procedure section of PI
| User presses buttons on the instrument screen | Testing Personnel |  | Instrument cancels test if HOME button pressed (and verified), test will not be completed. HOME is only button on the screen at this point | Test Procedure section of PI
Interpreting Results | Misinterpretation of test results | Testing Personnel |  | Instrument provides Flu A and B positive, negative, invalid result on result screen. Optional printer is available. | Result Interpretation section of PI
Disposal of test components | User does not connect Test Base to Sample Receiver | Testing Personnel |  | Instrument screen instructions, and feedback from disposables | Test Procedure section of PI, Precautions section of PI and Technical Memorandum Amplicon Contamination Prevention
| Test Base not properly connected to Sample Receiver | Testing Personnel |  | Instrument screen instructions, feedback from disposables (Test Base/Transfer Cartridge clicks into Sample Receiver). | Test Procedure section of PI, Precautions section of PI and Technical Memorandum Amplicon Contamination Prevention
External Control failure | Improper control storage | Environment |  | None | Storage and Stability section of PI
| Heads of positive and negative control swabs touched | Testing Personnel |  | None | Precautions section of PI
| Patient sample run as a control | Testing Personnel |  | Instrument screen has clearly marked button for positive and negative controls | Running a QC Test section of UM and PI
| Testing Positive control swab as Negative control and vice versa | Testing Personnel |  | Instrument screen has clearly marked button for positive and negative controls. Swab packages are clearly labeled as positive or negative. | Swab packages are clearly labeled as positive or negative.
Carryover | Instrument contamination | Instrument system |  | Closed system designed not to leak if handled properly | Precautions section of PI and Technical Memorandum Amplicon Contamination Prevention
Maintenance failure | User does not clean instrument | Testing Personnel |  | None | Introduction/Cleaning and Maintenance section of UM
| Cleaning materials stick to the instrument and cover the optics for the barcode | Testing Personnel |  | If barcode is not recognized because optics are covered an error message is given - Error: Barcode Not Recognized | Introduction/Cleaning and Maintenance section of UM
| Cleaning with non- approved reagents | Testing Personnel |  | None | Introduction/Cleaning and Maintenance section of UM
Data memory full | User doesn't export results | Testing Personnel |  | Warning screen shown when memory almost full | Review Memory section of UM
Data loss | User doesn't export results or user lists before software upgrade | Testing Personnel |  | Data can be backed up on USB. | Review Memory section of UM
Software upgrade failure | Wrong USB drive used for software upgrade | Testing Personnel |  | Instrument screen instructions. Error: USB drive not found The instrument has not recognized the USB drive. | Software Updates section of UM
| Software update not upgraded | Testing Personnel |  | None | Software Updates section of UM
| USB removed during software upgrade | Testing Personnel |  | None | Software Release 5.2.2 -pro USB Software Update section
| Software Update not completed correctly | Testing Personnel |  | Instrument performs a system check and instrument screen requires QC Testing after software update. | Software Updates section of UM
© 2019 Abbott. All rights reserved. All trademarks referenced are trademarks of either the Abbott group of companies or their respective owners. 120004682 v03 04/19
Note on G7: Smith, Jane L: